## Slide 1: An S-Band Interlaced Side-Coupled Standing-Wave Linac with Continuous Energy Tunability from 1 to 10 MeV for Radiotherapy
Haipeng Zhi
Jiaru Shi, Hao Zha, Xiancai Lin, Jiaqi Qiu, Huaibi Chen
Department of Engineering Physics, Tsinghua University
Key Laboratory of Particle and Radiation Imaging, Ministry of Education With University of Electronic Science and Technology of China
OPAC2026

## Slide 2: Medical motivation for 1–10 MeV tuning
02
Different medical applications motivate broad-range electron-energy flexibility.
This work: continuous 1–10 MeV source (design target)
1
2.5
9
Electron energy (MeV)
10 12
Portal X-ray imaging [1]
Electrons converted at a target
4 6
Incident 2.5 MeV electrons → Cu target
Dedicated 2.5 MV bremsstrahlung imaging beam
Electron radiotherapy [2]
Depth selection for superficial targets
Nominal clinical range: 4–22 MeV
continues beyond this axis
Electron IORT [3]
Direct irradiation during surgery
Mobetron IORT: 6, 9 and 12 MeV presets
Imaging needs
Electron energy and target design determine the X-ray spectrum and influence image contrast, penetration and dose.
Treatment needs
Electron energy governs therapeutic penetration and distal dose falloff, enabling
depth-matched treatment.
Complementary imaging and therapy applications motivate broad-range energy flexibility in compact electron linacs.
[1] Gräfe et al., JACMP 17(5):222–234 (2016). doi:10.1120/jacmp.v17i5.6323. [2] IAEA, Radiation Oncology Physics (2005), Ch. 8.
[3] IntraOp, Mobetron HDR Product Datasheet, Rev. D. Clinical presets are nominal energies; source and patient-surface energies can differ.
Representative electron energies
Representative applications;
energy ranges are not strict boundaries.

## Slide 3: Low-MeV electrons for low-MV imaging
03
A clinical example: TrueBeam portal imaging uses 2.5 MeV electrons incident on a copper target [1].
Clinical platform Electron-to-photon conversion [1]
2.5 MeV electrons 2 mm Cu target 2.5 MV X rays
Bremsstrahlung has a broad photon spectrum. The MV label does not mean that every photon has an energy of 2.5 MeV.
Published image-quality comparison
MV high-quality mode; DMI detector. Gräfe et al., Fig. 5 [1], CC BY 3.0.
The 2.5 MV imaging beam shows improved low-contrast visibility compared with the 6 MV beam
under the reported phantom conditions.
Gräfe et al., “Characterization of a 2.5 MV inline portal imaging beam,” JACMP 17(5):222–234 (2016). doi:10.1120/jacmp.v17i5.6323.
Varian / Siemens Healthineers, TrueBeam official product page. The phantom comparison establishes an advantage over 6 MV, not general superiority over kV.
TrueBeam — Varian
Manufacturer image [2] Representative platform illustration
Leeds TOR 18FG phantom, without overlying Solid Water. Low-contrast targets visible [1]
18/18 at 2.5 MV
10/18 at 6 MV

## Slide 4: Electron radiotherapy and depth matching
04
Gao et al., Fig. 1 [1], CC BY 4.0. 10 × 10 cm² cone, SSD = 100 cm.
The 11/12 MeV beams provide clinical context beyond this design’s 10 MeV limit.
Clinical energy presets control the distal dose falloff. Intermediate energies provide finer depth selection [1,2].
Direct electrons: rapid distal dose falloff [2] Clinical energy-selection platforms
Versa HD · Elekta
Photon/electron linac Clinical presets [3]
Mobetron · IntraOp
Electron IORT 6, 9, 12 MeV [4]
Custom 7 and 11 MeV settings [1]
Electron setting | Target I₅₀
7 MeV | 2.9 cm
11 MeV | 4.2 cm
Target I₅₀: 50% ionization depth.
15 × 15 cm² cone, SSD = 100 cm.
≈0.5 cm R₈₀ shift for 6→7 and 11→12 MeV
(10 × 10 cm² cone, SSD = 100 cm) [1].
Clinical demand for finer depth matching motivates continuous 1–10 MeV source-energy control.
[1] Gao et al., JACMP 23:e13633 (2022). doi:10.1002/acm2.13633. [2] IAEA, Radiation Oncology Physics (2005), Ch. 8.
[3] Elekta, Versa HD official product and commissioning materials. [4] IntraOp, Mobetron HDR Product Datasheet, Rev. D. Device images: manufacturers.

## Slide 5: Conventional energy-tuning methods
05
Independently powered sections
Side-cavity energy switching
How can net energy gain be reduced without sacrificing strong RF focusing?
Proposed: Interlaced Dual-Chain Linac
Key challenge: Reducing the net energy gain without weakening local RF focusing.
Principles and Limitations
01 | RF Power Modulation
Reduced bunching fields → lower capture efficiency and increased energy spread.
02 | Independently Powered Sections
Independent RF amplitude/phase control → increased system complexity.
03 | Side-Cavity Energy Switching
Downstream field adjustment → discrete energy selection.

## Slide 6: Interlaced dual-chain architecture
06
2 bunching+5 accelerating cells 6 accelerating cells
COMBINED INTERLACED STRUCTURE
Functional separation + electromagnetic independence enable voltage-superposition tuning.
Φ6 mm iris
A
B
Electron
DUAL-CHAIN ARCHITECTURE
CHAIN A
Fixed RF: bunching and capture
CHAIN B
Tunable RF: energy control
350 mm × Φ150mm
13 on-axis cells I 2998MHz
Two independent RF feeds
MICROWAVE ISOLATION
Adjacent A/B cells are connected only through the beam iris
RF operating freq.
2.998GHz
TE11 iris cutoff
~29GHz
Inter-chain coupling.
＜0.01%
Intra-chain coupling.
0.8%
Strong RF isolation enables independent amplitude and phase control.

## Slide 7: RF design and phase synchronization
07
2998 MHz · π/2 side-coupled mode
Rising bunching field
Uniform accelerating field
A
B
A
90° flight 90° flight
180° transit between A cells compensates their 180° field reversal
Regular cell | Value
Length | 25.0 mm
Transit-time factor | 0.952
ZT² | 99.25 MΩ/m
Quarter-wavelength cells maintain synchronism within each interlaced chain.
𝜆Τ4 single cell

## Slide 8: Two tuning modes and the transition
08
High energy
5 < E ≤ 10 MeV
B
A
B
A
B
A
B
+ + + + + + +
Transition
Nominal E = 5 MeV
B
A
B
A
B
A
B
0 + 0 + 0 + 0
A accelerates
B is off
Chain B voltage: zero
Low energy
1 ≤ E < 5 MeV
B
A
B
A
B
A
B
− + − + − + −
A accelerates
B decelerates
Δφ = +90°
Reversing Chain B by 180° reduces net gain while preserving strong local fields.
Definition: Δφ =φB- φA
A accelerates
B accelerates
Δφ = −90°

## Slide 9: Circuit model and operating points
09
Resonant power balance with beam loading
Circuit quantity | 10 MeV mode | 6 MeV mode | 1 MeV mode
Chain A power (MW) | 2.10 | 2.10 | 2.10
Chain B power (MW) | 2.10 | 0.15 | 1.00
Chain A voltage (MV) | 5.45 | 5.45 | 5.45
Chain B voltage (MV) | 4.90 | 0.84 | −4.48
Relative phase | −90° | −90° | +90°
Total voltage (MV) | 10.35 | 6.29 | 0.97
The RF power determines the voltage branches, while the relative phase selects acceleration or deceleration.
V +
Accelerating solution
Energy transferred from RF field to beam.
V −
Decelerating solution
Energy transferred from beam to RF field.
I = 100 mA ; βc = 1.30 R_A / R_B = 19.1 / 14.9 MΩ

## Slide 10: Continuous tuning by voltage superposition
10
High-energy mode
5 < E ≤ 10 MeV
Δφ = −90°
Transition
E ≈ 5 MeV
Chain B off
Low-energy mode
1 ≤ E < 5 MeV
Δφ = +90°
Amplitude controls voltage magnitude; phase reversal selects acceleration or deceleration.

## Slide 11: Comparison of Four Energy-Tuning Methods
11
D Energy switching
Same 11-cell downstream section and identical post-buncher distribution
1.2 MeV average energy, 5.4% rms energy spread, σx = 1.1 mm.
A Alternating accel./decel. B Uniform RF reduction
Axial field (MV/m)
C Sectioned phase control
Particle energy (MeV)
Equal 3 MeV exit energy provides a common basis for comparing transverse transport.
Sec. III.A, Fig. 5(a,b)

## Slide 12: Single-Particle Transverse Dynamics Model
Both accelerating and decelerating cells can provide net RF focusing
due to energy-dependent field weighting.
Sec. III.B.1, Eqs. (4–6); β = v/c, prime = d/dz, e > 0
12
Linearized paraxial motion
𝑥″ +
𝛾′
2
𝛽 𝛾
𝑥′ + 𝑘
2
RF
𝑧 − 𝑘sc
2
𝑧 𝑥 = 0
(4)
RF restoring coefficient
RF
𝑘2
𝑧 =
𝑒
2𝑚𝑐2𝛽2𝛾
𝜕𝐸
𝜕𝑧
𝑧 +
𝛽 𝜕𝐸
𝑧
𝑐 𝜕𝑡
(5)
Uniform cylindrical bunch
sc
𝑘2
𝑧 =
𝐾
𝑥
𝛽2𝛾3𝜎2
(6)
𝐾 =
𝑒𝐼
0
2𝜋𝜀 𝑚𝑐3
Local mechanisms
γ′ > 0: adiabatic damping
γ′ < 0: antidamping
RF focusing depends on the local
field and particle energy.
Lower beam energy strengthens space-charge defocusing.
Ez = Ea(z) sin(ωt + φ)
σx: rms size. K: generalized perveance.
Accelerating cell: entrance half focuses. Decelerating cell: exit half focuses.
Lower local energy gives these intervals greater weight, with |k²RF| ∝ 1/(β²γ).
(7)

## Slide 13: RF Focusing versus Space-Charge Defocusing
Method A maintains stronger transverse focusing despite stronger space-charge defocusing.
Sec. III.B.1, Fig. 6(a,b); k²RF is a signed coefficient
13
RF focusing coefficient Space-charge defocusing coefficient
A: lower mean energy, stronger space charge. C: higher mean energy, weaker space charge.
A, C: strong local RF fields and focusing. B, D: substantially weaker RF focusing.
Typical |k²RF| is about 10² times k²sc at this operating point.

## Slide 14: Cumulative Focusing and Particle Trajectories
The alternating method provides the strongest cumulative focusing
and superior confinement of the probe particle.
Sec. III.B.1, Eqs. (7–9), Fig. 6(c,d)
14
I = 100 mA, σx,0 = 1.1 mm (fixed) x0 = σx,0, x′0 = 0
A has the largest cumulative focusing.
C: sustained deceleration causes antidamping.
B, D: weaker RF fields and progressive divergence. Probe trajectories support the cumulative comparison.
Cumulative focusing Φeff (m⁻¹)
RF force dominates space charge in this reference
calculation.
Probe-particle displacement x (mm)

## Slide 15: Multi-Particle Validation with ASTRA
Self-consistent multi-particle tracking confirms the transverse-focusing advantage
of the alternating acceleration–deceleration method.
Sec. III.A and III.B.2, Fig. 7(a,b)
15
A: bounded after initial oscillations. B, D: growth. C: envelope growth in the consecutive decelerating cells.
A 29.1% B 21.2% C 25.8% D 24.7%
A versus B: +7.9 percentage points
Capture includes bunching losses. Nonlinear space charge and finite aperture affect the results.
ASTRA: 10⁵ macroparticles with self-consistent space charge. Reference current: 100 mA.
Same post-buncher distribution, same 11-cell section and identical final energy of 3.0 MeV.
A Alternating accel./decel. B Uniform RF reduction C Sectioned phase control D Energy switching

## Slide 16: Beam Quality across the 1–10 MeV Tuning Range
Beam quality varies smoothly across 1–10 MeV, with no abrupt deterioration at the 5 MeV mode transition.
Sec. III.C, Fig. 8 and Table IV; 0.5 MeV scan with synchronous-phase optimization
16
(a)
(b)
(c)
(d)
Capture above 29% throughout
Overall capture increases
with output energy.
Smaller transverse rms size
0.89 mm at 1 MeV to 0.60 mm at 10 MeV. Consistent with weaker space-charge defocusing.
Lower energy spread near 6 MeV
Strong decrease from low energy,
followed by a slight rise at high energy.
Shallow U-shaped bunch length
Longitudinal phase focusing shifts the waist relative to the structure exit.
5 MeV: Chain B changes phase
Purple dashed lines mark the transition. No abrupt beam-quality degradation.

## Slide 17: Representative Energy Spectra and Transverse Phase Space
The 1 MeV mode shows stronger nonlinear distortion. The 6 and 10 MeV modes maintain improved transverse beam quality.
Sec. III.C, Fig. 9 and Table IV; 100 mA pulse current; nominal modes and mean energies distinguished
17
Mean E: 1.06 MeV Rms x′: 2.01 mrad Nonlinear space charge and larger size increase nonlinear RF-field sampling.
1 MeV mode 6 MeV mode
Mean E: 6.05 MeV Rms x′: 1.29 mrad Less distortion, with the smallest divergence. Favorable energy spread and bunch length.
10 MeV mode
Mean E: 10.03 MeV Rms x′: 1.36 mrad Weaker space charge improves confinement. Waist proximity raises divergence above 6 MeV.

## Slide 18: RF implementation and tuning accuracy
18
Two 2.5 MW magnetrons with master-slave injection locking
1° phase precision
<0.1% energy contribution
0.1 dB attenuation resolution
Up to 5.2% at 1 MeV before calibration
Near-cancellation makes low-energy accuracy most sensitive to Chain B amplitude control.
Manuscript Figs. 10–11, Eqs. (10–11); proposed system, supported by cited magnetron-locking experiments

## Slide 19: Cooling and differential thermal detuning
19
Inlet water: 20 °C ; Mean velocity: 2 m/s ; Flow rate: 38 L/min Symmetric operation Asymmetric operation Embedded serpentine channels Both chains at 2.1 MW Chain A on, Chain B off |  |  |  | 
 | Thermal result | Symmetric | Asymmetric | 
 | Peak temperature | 38.4 °C | 29.3 °C | 
 | Frequency shift A / B | −219 / −205 kHz | −84 / −76 kHz | 
 | Differential shift magnitude | 14 kHz | 8 kHz | 
Differential thermal detuning stays below 20 kHz in both simulated operating cases.

## Slide 20: 1–10 MeV
Conclusions and Outlook
Continuous Energy Tuning
Independent chain amplitude and phase
29.1%
Capture Ratio at 3 MeV
Highest simulated capture of four methods
<20 kHz
Inter-Chain Frequency Difference
Thermal simulations, both operating cases
01
02
03
Continuous Energy Tunability
Independent chain voltages tune the output energy while Chain A keeps the bunching conditions fixed.
Enhanced Low-Energy RF Focusing
Alternating acceleration and deceleration maintains
strong RF fields, improving confinement and simulated capture.
Engineering Feasibility
Phase-locked dual-magnetron RF and cooling designs support independent RF control and thermal stability.
Compact Interlaced Linac
Accelerating structure: 350 mm length × 150 mm diameter
Prototype Fabrication
→
→
Beam Measurements
RF Commissioning
Thank you for your attention.
Sec. V and Fig. 14; design and simulation results. Prototype validation is the next stage.
20